## Slide 1
Notes: The authors of this presentation from the Cleveland Clinic, University of Pittsburgh and Sportsguard Laboratories, would like to thank the FAA Triennial Conference Program Committee, and in particular, the reviewers, for allowing our work to be included in this year’s Conference.
Today’s presentation summarizes the authors’ work toward “Determining Loss of Consciousness in Fighters and Development of Associated Injury Assessment Reference Values.” [CLICK]

## Slide 2: CONFLICT OF INTEREST DISCLOSURE
Cleveland Clinic entered into an exclusive license agreement with Prevent Biometrics for Intelligent Mouthguard technology on which the Authors are inventors. As inventors, they are entitled to a portion of any commercialization revenues Cleveland Clinic receives. The Authors are also paid consultants of Prevent Biometrics and members of Prevent Biometrics’ Scientific Advisory Panel.
Notes: Due to potential conflict of interest, the audience is encouraged to consider adding extra scrutiny to the information presented, and be compelled only by the objective data presented. [click]

## Slide 3: Overview
INTRODUCTION – injury risk and measurement uncertainty
BACKGROUND – mouthguard and rigid body assumption
METHODS – validate the instrument in the laboratory, collect human head impact data.
RESULTS – laboratory validation and human head impact data
DISCUSSION
CONCLUSIONS
ACKNOWLEDGEMENTS
Notes: In the introduction, the authors describe the relationship between the injury risk and the single head impact dose measurement uncertainty.
Basic assumptions behind the instrument theory are described in the Background section.
Laboratory validation and human head impact data collection are described in the Methods.
The Results show comparison of the laboratory validation of different head impact dosimeters and the human head impact data.
The Discussion and Conclusions show comparison of previously collected human head impact data vs. the Cleveland Clinic data, along with further implementation plans.
[CLICK]

## Slide 4: INTRODUCTION – injury risk and measurement uncertainty

## Slide 5: Introduction
US: ~173,000 athlete concussions yearly <19 years old
  Short-term: Diffuse Axonal Injury, Subdural Hematoma, Transient Neurological Dysfunction
  Long-term: Alzheimer's Disease, Parkinson's Disease, Chronic Traumatic Encephalopathy, Depression
Short-term, long-term consequences function of:
  Physiology (patient & population specific)
  Impact doses (head impact location, duration)
  Dose response (single, cumulative)
Our long-term goal is to develop concussion risk functions for frontline caregivers
Our short-term goal is to create and deploy a calibrated single event head impact dosimeter (“Intelligent Mouthguard”)
Notes: 5
5
Concussion is a major problem for young athletes in the United States and causes both short-term and long-term consequences. [CLICK]
Consequences depend upon physiology, impact doses (e.g. collision severity) & response to impact doses. [CLICK]
Our long term goal is to develop risk functions for caregivers; our short term goal is to create and deploy a calibrated single event dosimeter we call the “Intelligent Mouthguard” [CLICK]

## Slide 6
Chart: Dose (units)
 | 0 | 0.5 | 0.5 | 0.5
Smaller Dose Uncertainty
Smaller Risk Uncertainty
Larger Dose Uncertainty
Larger Risk Uncertainty
Notes: Risk uncertainty in identifying a dangerous head impact dose is a function of single event dose measurement uncertainty. [CLICK]
Assuming this s-shaped red curve represents our desired risk function, it becomes obvious that to minimize uncertainty in the risk function, one needs to minimize uncertainty in dose measurement. [CLICK]
Said differently, frontline caregivers require the lowest practical injury risk uncertainty; lower measurement uncertainty helps achieve this. [CLICK]

## Slide 7: BACKGROUND – mouthguard and rigid body assumption
Notes: An instrumented mouthguard provides the best practical “sensor to skull” coupling, and is capable of characterizing the movement of the head as a rigid body (not skull deformations).

## Slide 8: Head Impact Dosimetry Origins
1947-1967 research on head acceleration (mostly on himself)
421 mph, 46g
Edwards AFB Archives
Col. JP Stapp, MD PhD
Notes: Head impact dosimetry can trace its roots backwards at least 70 years, with Col. John Paul Stapp being a key pioneer in this field.
The original “human crash test dummy”, Col. Stapp’s self-experiments are notable for a case where, when strapped to a jet-powered sled, he decelerated from 421mph to a dead stop.
In this test, his eyeballs nearly popped out of his skull – a vigorous test of human tolerance to deceleration injury! [click]

## Slide 9: Origins of Head Impact Dosimeters
“Col. John Paul Stapp recommended instrumenting football players … as long ago as 1975…”
-David Foust, 1978 Int’l Biomechanics Workshop
“…is it possible to put (sensors) in the mouthpiece? …that’s the best coupling you can ever get.”
-Dr. John Melvin, 2003 Int’l Biomechanics Workshop
Notes: In addition to Col. Stapp in the 1970’s, John Melvin and others in the DoD, DOT and academia taught us that instrumented mouthpieces could be used to study head impact dose and injury risk. [CLICK]
We certainly tip our caps to Dr. Melvin for pointing out in 2003 sensor-skull coupling in a mouthpiece is better than helmet or skin-mounting. [CLICK]

## Slide 10: Commercial Dosimeters
Brain Sentry
2003-Present
Head Impact Telemetry System (HITS)
($1,000-$10,000)
2010-Present
First Alert
I1 Biometrics
Notes: There are many commercial dosimeters available, from the HITS system (on the left) to new and emerging technologies (on the right).
Many of these systems provide scalar outputs only, and few have been sufficiently calibrated in the laboratory to determine single head impact measurement uncertainty [click through slide]

## Slide 11: Intelligent Mouthguard (2008-Present)
2008-2010
3 Channels, demo
2011-2014
6 Channels, wired
2015-
12 Channels, wireless
Notes: 11
11
At Cleveland Clinic we have been working on measuring single head impacts for a number of years.
Early versions had 3 channels, suitable for 3-degree-of-freedom measurement. [click]
A later version had 6 channels but an extra-oral cable. [click]
The current version has 12 channels, and includes wireless charging and data offload. Each accelerometer has full scale 200g range, can store 64 events on-board for 50 milliseconds per event. The circuit board is pressure-lamitated between layers of EVA mouthguard material.
The theoretical rigid body solution to calculate Reference cg acceleration from mouthguard sensors requires us to quantify sensitive axis orientations and sensor locations for each user’s mouthguard. [CLICK]

## Slide 12: Rigid Body Assumption
Calculate REF Hybrid III accel () from Intelligent Mouthguard (IMG) measurements (, , ) and known geometry () for 6DOF spatial (location, direction) & temporal (duration) parameters
Three regions of interest (R1, R2, R3)
  We focus on R1 up to peak.
Notes: 12
12
We are using well-known rigid body principles to quantify six-degree of freedom head motion in response to impact based on Intelligent Mouthguard sensor data. [CLICK]
For convenience of analysis we divided each impact profile time trace into three regions: R1, R2 & R3. While in calibration tests we focused mainly on Intelligent Mouthguard output versus Reference Hybrid III during R1 up to time of peak, we also calculated mouthguard data versus reference in R2 and R3. PAUSE A FEW SECONDS…. [CLICK]

## Slide 13: METHODS – validate the instrument in the laboratory, collect human head impact data.
Notes: The authors’ approach is that one can trust the human impact data as long as the engineering parameters of the head impact fit within the laboratory validation range.

## Slide 14: In Vitro Test Method
Linear pneumatic impacts to Hybrid III “3-2-2-2 + 3w” headform
15kg mass, 2.0m/s to 8.5m/s
American football helmet (HITS Revo Speed, n=178) and boxing headgear (n=68)
Variety of impact locations, dentition-only coupling (jaw clamping F=0)
Linear acceleration peaks up to ~120g
Notes: In vitro, we used our linear impactor to strike a Hybrid III equipped with a reference “3-2-2-2 plus 3-omega” sensor package in a variety of impact configurations and impact velocities.
We applied no additional clamping force to the Intelligent Mouthguard, it was held on the dentition model by friction only.
Reference Hybrid III linear CG acceleration peaks up to approximately 120g were tested. [CLICK]

## Slide 15: Human Impact Dose Collection
Amateur boxers (n=4, 15-18 YO) and American football (n=8, 11-20 YO)*
  Boxing: sparring, five, 3-minute rounds
  American football: four separate practices and scrimmages/games
  Report on confirmed true positives
    We confirmed true positives with video and algorithmic data analysis.
*Cleveland Clinic IRB 13-899
Notes: In our human tests, we collected single event head impact doses on American football players during practice and scrimmages, as well as amateur boxers during competitive sparring. [CLICK]
Here I am reporting on true positive head impact doses, confirmed by our video analyses. [CLICK]

## Slide 16: RESULTS – laboratory validation and human head impact data
Notes: Cleveland Clinic mouthguard has superior laboratory validation results, and has been used as a single event head impact dosimeter in the field.

## Slide 17
Chart: Scalar Acceleration Laboratory Calibration Comparison
| 0 | 20000
 | 0 | 20000
Notes: 17
17
In the laboratory, mouthguard acceleration outputs versus Reference Hybrid III fit a linear model of the form Y=1 * X + 0g at time of peak. While scalar outputs are shown here, XYZ components also fit a linear model.
As a comparative analysis, calibration data collected by Siegmund et al. (2015) for the National Football League on an instrumented helmet and different mouthguard are presented on the same plot.
Note the relative scatter in the comparative datasets; neither instrument tested by Siegmund et al. fit a linear model versus reference. [click]

## Slide 18: Youth Football Dose
Direction, Location and Magnitude for all Impacts in Boxing, American Football
All Impact Doses
11-yr old head impact
Notes: In this slide, you will see a single event head impact dose received by an 11-year old American footballer. The first harmonic in this impact was 56Hz, within the range of laboratory calibration impacts. [CLICK]
The figure on bottom left shows the time-varying X, Y, and Z-accelerations at the player’s center of gravity after using our coupling, custom data and instrument linearization techniques.
The figure on right left shows this impact, along with all other impacts in boxing and American football. [CLICK]

## Slide 19: Impact Dose Results Compared w/HITS
We had no concussions or clinical symptoms, yet many Intelligent Mouthguard head impacts near concussive ranges collected by HITS Helmets
Notes: When compared by the scalar peak values of the linear acceleration and the angular acceleration, the Cleveland Clinic non-concussion human head impact data sets show overlap with the published concussion data sets. Among the possible explanations may be:
variation in individual tolerance to the head impact , and
inadequacy of the scalar acceleration peak metric.

## Slide 20: CONCLUSIONS
There is an unmet need for an accurate and precise head impact dosimeter to help identify those at-risk
  Intelligent Mouthguard measures 6-degree of freedom impact kinematics as time-dependent vectors and scalars
  Laboratory data confirm Intelligent Mouthguard is only known dosimeter to meet NFL validity specification
  Now possible to monitor vector-based single impacts in real-time, as opposed to analyzing scalar impact distributions
Compared to HITS scalar concussion doses, non-concussion Intelligent Mouthguard doses overlap
  Scalar comparisons obscure direction/location differences
  Physiological variability among athletes may also play a role
Notes: The authors are working on implementation of a laboratory validated single event head impact dosimeter, which will provide datasets necessary for determining loss of consciousness parameters in fighters and the development of injury assessment reference values.

## Slide 21: FUTURE WORK
Currently deploying Intelligent Mouthguard to American football, boxing and MMA subjects in Las Vegas, Minneapolis, Dallas and Milwaukee
  Approx. 100 athletes currently under study
  Male and female
  Ages 12-22
Data: ‘false positives’ and ‘true positives’
  ‘false positives’ = recorded non-head impact events: filtered out through signal analysis
  ‘true positives’= recorded and confirmed head impact events : contribute to head impact injury assessment reference values (IARV)
Plans to expand studies in 2017 to rugby & soccer
Notes: [read verbatim]

## Slide 22: Acknowledgments
PEOPLE:
Basar Aka, Anno van den Akker, Alper Aksu, Jay Alberts, Erik Beall, Gordon Bell, Ed Benzel, Dan Brett, Tara Bonner, Dave Conger, Robb Colbrunn, Lars Gilbertson, Frank Grazzini, Paul Haynes, Erica Heisler, Joshua Hirsch, Steve Jones, Barry Kuban,Mark Lowe, Prasath Mageswaran, Gil Martinez, Vince Miele, Mike Modic, Caglar Ozturk, Brian Perse, Brach Poston, Vikas Prakash, Trent Stratton, Kevin Waters, Steve Washburn, Karl West, Michael Wright
FUNDING:
Cleveland Clinic Innovations
Cleveland Clinic Neurological Institute
Cleveland Clinic RPC
Federal Aviation Administration (15-G-012)
NFL Charities
NIH R41HD079158
Orthopaedic Research and Education Fund
Prevent Biometrics
Rawlings Sporting Goods
Sportsguard Laboratories
Zielony Research and Education Fund

## Slide 23: THANK YOU