## Slide 2
Motivation
Gas phase active flame retardants and mineral fillers are broadly applicable (as opposed to char promoting compounds).
Halogens (especially Br) are effective at low loading and are the only commercial materials exhibiting flame inhibition, but they are under duress.
Are there other atoms/compounds that are gas phase active?
If so, how can we find them, and what is their combustion toxicity?

## Slide 3: Scales of Fire Behavior & Testing
???
No methods available for studying fire behavior at the milligram (lab) size.

## Slide 4
Approach
Develop rapid laboratory test to screen flame retardants and FR plastics for:
  Gas phase activity
  Combustion toxicity
103 kg
1 kg
10-6 kg

## Slide 5
Strategy
  Examine gas phase combustion chemistry using Microscale Combustion Calorimeter (MCC) to control
    Combustion temperature (Safronava, BCC 2016)
    Fuel/oxygen () ratio
  Measure combustion gases using
    CO and CO2 analyzers
    Fourier Transform Infrared (FTIR) spectrometer
Measure Fire Effluents at Micro Scale

## Slide 6
Background: Full Scale Fire Tests
Generation of Fire Effluents at Full or Large Scale
ISO 9705 /ASTM E 2257
Room Corner Test
ISO 24473
Furniture Calorimeter
EN 1323
Single Burning Item (SBI)

## Slide 7
Background: Bench Scale Fire Tests
Vitiated Fire Effluents at Bench Scale
Controlled Atmosphere
Cone Calorimeter
Fire Propagation
Apparatus
ASTM E 2058
Purser (Steady State Tube) Furnace
IEC 60695-7-50
FTIR

## Slide 8: Fuel / Oxygen Ratio in Full & Bench Scale
GER < 1 = Fuel Lean (over-ventilated)
GER > 1 = Fuel Rich (under-ventilated)
GER
=
Mass Loss Rate (kg-fuel/s)/Ventilation Rate (kg-O2/s)
(mFuel/mO2 )stoichiometric
GER is the global equivalence ratio in fire tests

## Slide 9
Combustion Chemistry
Fuel + Some Oxygen  CO + HCN + Hydrocarbons + Other Toxics
Fires generate toxic combustion products when oxygen becomes depleted (vitiated),
Incomplete & toxic combustion products
Complete (less-toxic) combustion products
Excess Oxygen
Insufficient Oxygen
CO2 + H2O + N2
Late Stage Fire
Early Stage Fire

## Slide 10
Toxic Gas Yields in Vitiated Full/Bench Scale
0.1
1
10
100
0
0.5
1
1.5
2
2.5
HCN Yield (g/kg-volatiles)
0.1
1
10
100
1000
0
0.5
1
1.5
2
2.5
ISO Room
ISO Room
(1/3-Scale)
Controlled Atmosphere
Cone Calorimeter
CO Yield (g/kg-volatiles)
Global Equivalence Ratio, GER
Global Equivalence Ratio, GER
Controlled
Atmosphere
Cone Calorimeter
FPA
ISO Room
Polyhexamethyleneadipamide / PA66
Carbon
Monoxide
Hydrogen
Cyanide
(A. Stec & R. Hull, Fire Toxicity, 2010)

## Slide 11: Combustion Process of Solid Polymers
* J. Troitzsch, International Plastics Flammability Handbook, 2nd Edition, Chapter 4, Hanser Publishers, Munich Germany, 1990
Combustion
Pyrolysis
Exothermic, -∆
Combustion Products
+ O2
Thermal Feedback from Flame
Combustible +
Non-combustible Gases
Liquid Products + Tar
Solid Charred Residue
POLYMER
Endothermic, +∆

## Slide 12
Microscale Combustion Calorimeter

## Slide 13
Heat Release Rate
Heat release rate by oxygen consumption normalized to initial sample mass
Peak rate of heat release
Heat Release Capacity
… dependent only on material properties

## Slide 14: Standard Test ASTM D 7309 (Method A)
Heat of Combustion, Q
(Fire Load)
0
100
200
300
400
500
600
700
450
500
550
600
650
700
Temperature (C)
Onset (Ignition)
Temperature
Peak Pyrolysis
Temperature
Heat Release Capacity, c
Pyrolysis Residue 
(Char Fraction)
Heat Release Rate Q/ (J/g-K)
Anaerobic pyrolysis at 1 K/s + complete combustion of gases at 900C, 20% O2
Microscale Combustion
Calorimeter

## Slide 15
Controlled Combustor Temperature
Fixed Flame Temperature
O2 + N2
(Air)
Fuel
Fuel
Plastic
Products
Products
Combustion in MCC versus in Flame
Premixed, Nonflaming
Combustion (MCC)
Diffusion-Limited
Combustion (Flame)
Premixed
Combustion
Fuel/O2/N2
O2
Controlled Fuel/O2 ratio
N2
Diffusion Controlled Fuel/O2

## Slide 16: Fuel / Oxygen Ratio in MCC
 < 1 = Fuel Lean (excess oxygen)
 > 1 = Fuel Rich (oxygen starved)
F
=
mFuel / mO2
 is the fuel/oxygen mass ratio relative to the stoichiometric mass ratio.
(mFuel/mO2 )stoichiometric

## Slide 17
 = 1 K/s
Specific Heat Release Rate, Q (W/g)
Sample Temperature (C)
0
200
400
600
800
1000
1200
0
0.1
0.2
0.3
0.4
0.5
0
50
100
150
200
qext = 50 kW/m2
HRR, kW/m2
Global Equivalence Ratio, GER
Elapsed Time, seconds
Local Equivalence Ratio, 
0
200
400
600
800
1000
0
0.1
0.2
0.3
0.4
0.5
350
400
450
500
550
 In Standard MCC vs GER in Standard Cone
MCC (ASTM D 7309)
Cone Calorimeter (ASTM E 1354)
High Impact Polystyrene/HIPS

## Slide 18: Fuel / Oxygen Ratio in Flame
Fuel
Air

0
∞
Distance into Flame Sheet
Varies from ∞ to 0 in diffusion flame
flame = Average Value
Fuel
Gases
Laminar
Diffusion
Flame
Sheet
Air
Combustible Solid

## Slide 19: Method of Controlling  in MCC
0
0.5
1
1.5
2
2.5
3
200
300
400
500
600
700
Percent of Oxygen Consumed, ∆ XO2
Sample Temperature, T (C)
Training Set
Program ( = 1)
Measured ( = 1)
Run material of mass m0 at  = 1 K/s, XO2 = 0.2, and Tc = 900C.
Measure ∆XO2/m0 at each sample temperature T for complete combustion of gases to obtain training set, (∆XO2/m,T)0.
Store training set (XO2,T,m)0 for material.
Use (∆XO2/m,T)0 data to control O2/N2 flows such that:
(∆XO2(T)/m)0
∆XO2(T)/m

=
=
m0
m
∆XO2(T)
(∆XO2(T))0

## Slide 20: Fuel / O2 Control and Measurement
Pyrolysis
Pyrolyzer
Cold Trap
Flow
Meter
O2
Analyzer
O2
N2
Sample
Exhaust
Premixed
Combustion
CO2
CO
Control From Look-up Table

1000C
DAQ

## Slide 21: Combustion Gas Analysis
Mass flow rate of gas
Where i = O2, CO, CO2
Time integral for total amount of gas consumed/produced
Yield of gas per mass of volatile fuel in sample

## Slide 22: Maximum Gas Yields
Polypropylene
1 C3H6 + 4.5 O2 3 CO2 + 3 H2O
Stoichiometry

## Slide 23
Mass Balance on Product Yields
Charring
Non-Charring

## Slide 24: MCC Combustion Gas Histories
0
1
2
3
4
5
6
7
8
Excess O2 Volume Fraction, %
Sample Temperature, C
300
350
400
450
500
Oxygen
(O2)
4
CO Volume Fraction, %
0
1
2
3
Carbon
Monoxide
(CO)
Sample Temperature, C
1
2
3
4
5
6
7
0
300
350
400
450
500
CO2 Volume Fraction, %
Carbon
Dioxide
(CO2)
2.0
1.5
1.0
0.75
0.5

Polymethylmethacrylate
(PMMA)
Asymmetry suggests Fuel/O2 not perfectly synchronized or r0 changes with T
2.0
1.5
1.0
0.75
0.5

2.0
1.5
1.0
0.75
0.5


## Slide 25
0
1
2
3
4
5
6
0
0.5
1
1.5
2
2.5
PE
HIPS
ABS
PA66
PMMA
POM
CO/CO2 Ratio (g/g)
Fuel/Oxygen Ratio (Relative to Stoichiometric)
Combustion Chemistry Changes with Fire Stage
Fuel + O2  CO + H2O + 
CO + O2  CO2 + O

## Slide 26
0.001
0.01
0.1
1
10
0
0.5
1
1.5
2
POM
COx Yields kg/kg
Effective Equivalence Ratio, 
CO2
CO
0.01
0.1
1
10
0
0.5
1
1.5
2
PE
Effective Equivalence Ratio, 
CO2
CO
0.001
0.01
0.1
1
10
0
0.5
1
1.5
2
PMMA
Effective Equivalence Ratio, 
CO2
CO
0.001
0.01
0.1
1
10
0
0.5
1
1.5
2
PA66
Effective Equivalence Ratio, 
Premixed (MCC)
Diffusion Flame (Cone)
CO2
CO
COx Yields kg/kg
For Low Soot Polymers, GER  flame < 1
Matching Yields for CO & CO2 Gives flame in Cone

## Slide 27
0.01
0.1
1
10
0
0.5
1
1.5
2
HIPS
0.01
0.1
1
10
0
0.5
1
1.5
2
ABS
COx Yields kg/kg
Effective Equivalence Ratio, 
COx Yields kg/kg
Effective Equivalence Ratio, 
CO2
CO
CO2
CO
Matching Yields for CO & CO2 Gives flame in Cone
0.01
0.1
1
10
0
0.5
1
1.5
2
PVDF
COx Yields kg/kg
Effective Equivalence Ratio, 
CO2
CO
Premixed (MCC)
Diffusion Flame (Cone)
For Soot-Forming Polymers,
GER  flame  1

## Slide 28
1
10
100
1000
0
0.5
1
1.5
2
2.5
CO Yield (g/kg)
0.01
0.1
1
10
100
0
0.5
1
1.5
2
2.5
HCN Yield (g/kg)
Equivalence Ratio
Equivalence Ratio
Carbon Monoxide
Hydrogen Cyanide
Comparison of All Vitiated PA66 Yields
Fire Propagation
Apparatus
Purser Furnace
1/3 ISO Room
ISO Room
Controlled Atm.
Cone Calorimeter
MCC-FTIR
MCC-CO analyzer
(A. Stec & R. Hull, Fire Toxicity, 2010)

## Slide 29: Combustion Chemistry at c
POM
PMMA
PA66
PP
HDPE
PET
HIPS
DER 332
PC
ABS
PEI
PC 121R
PC ABS
PS
PVDF
PS FR
PPS
10% BBA
PVC
PC ABS FR
20% 542
20% BBA
30% BBA
40% 542
PH PC 35% F75
40% BBA
60% 542
PH PC 60% F286
80% 542
PH F297

## Slide 30: High Temp Combustion Furnace
Heater capable of reaching 1700°C
- Approaches flame temperatures
- Recreate flame chemistry

## Slide 31: MCC w/ FTIR Evolved Gas Analysis

## Slide 32: Present / Future Work
Generate new FT-IR calibration spectra & develop method more
Test more materials & flame retardants
Examine higher combustion temperatures
Derive a toxicity index

## Slide 33: Conclusions
MCC-FTIR technique shows promise for monitoring combustion gas species using milligram samples in a 15 minute test with accurate control of equivalence ratio.
The MCC-FTIR methodology will be used to screen novel chemical compounds (e.g., organometallics) for gas phase flame inhibition activity.
The toxic gas yields may be incorporated into mixture fraction models for CFD.