## Slide 1: Revised Rate of Heat Release Test Method October, 2016 Triennial Conference Atlantic City, NJ USA
Michael Burns, FAA Technical Center

## Slide 2: AGENDA
HRR Test Method
Background for Revision
Improvements
Prototype Data
Next

## Slide 3: Heat Release Rate Test Method
Required by Federal Aviation Administration
Qualify large surface area interior materials for use in commercial passenger aircraft
  Ceilings
  Stow bins
  Sidewalls
Federal Register 14 CFR parts 25 and 121
  Description
  Use
  Calibration
  Pass / Fail criteria

## Slide 4: Revised Heat Release Rate Test Method
Goal: Simplify and Streamline
Multi-laboratory studies in which the same material was tested at different locations indicate reproducibility of test results can be improved.
Many factors can contribute to poor agreement between test laboratories.
This presentation describes a detailed revision of the test apparatus conducted by the International Aircraft Materials Fire Test Working Group to better standardize components, procedures and calibration of equipment for this test method.

## Slide 5: Areas Of Improvements
  Airflow (Dual Flow to Single Flow)
  Standardized Construction
  Procedures
  Software
  Calibration
  Equipment
  Changes in Documentation (Chapter HR)

## Slide 6: Airflow – Single Flow Design
OSU: Orifice Meter, Natural 3:1 Airflow Split
HR2: Mass Flow Controlled

## Slide 7: Removal of Bypass Air and Components
  Chimney
  Baffle Plate
  Outer Pyramid
  Bypass Air Manifold
  Piping
  Surplus Gasket Material and Hardware
  Orifice Meter/Mercury Manometer

## Slide 8: Removal of Bypass Air and Components
PROS
  Single air source
    1 MFC Req’d
    Smaller Compressed Air System Req’d
  Easier to clean, inspect, maintain
  Fewer Components
  Quieter system
  Thermocouples operating in a better performance range
  Lower Calibration Factor (Approx. 4.0 kW/m2/mV or 0.09 kW/mv)
CONS
  Higher Exhaust Gas Temperatures (approx. 500 degrees F)

## Slide 9: Airflow
Defined Air Quality
  “Fee of oil mist, water and foreign particles”
Airflow
  Flow rate controlled using a Mass Flow Controller to replace orifice meter
  Calibrated To STP 0º C at 760 mmHg
  Accuracy: 2% Full Scale
  Flow Rate: 20 ± 0.4 SCFM
  Inlet Air Temperature: 72.5 ± 2.5 °F

## Slide 10
¼ inch Aluminum | Unchanged | Holding Chamber Door
 | Unchanged | Air Distribution Plate
0.036 +/- 0.003 Alum. | Not Specified | Exhaust Stack Covering (New)
0.125 ± 0.005 | Not Specified | 1/8 inch Retaining Rod
0.048 ± 0.004 | (0.049 ± 0.002) | Main Body
 | Not Specified | Globar Pan
 | Not Specified | Holding Chamber
 | Not Specified | Second Stage Plate
 | Not Specified | Lower Air Plenum
Insulation - Thermal Conductivity Criteria
Overlap Criteria
Component Metal Thickness & Tolerances Table
Standardized Construction
New Previous Component

## Slide 11: Standardized Construction
0.042 ± 0.004 | (0.042 ± 0.002) | Diamond Mask
0.036 ± 0.003 | (0.030 ± 0.008) | Reflector Plate
 | 24-Gauge Steal | Sample Holder Plate
0.030 ± 0.003 | (0.031 ± 0.002) | Outer Pyramid (Removed)
 | Not Specified | Radiation Doors (2)
0.018 ± 0.002 | Unchanged | Inner Pyramid (Modified)
 | Unchanged | Chimney (Removed)
 | Unchanged | Baffle Plate (Removed)
 | Unchanged | Sample Holders
 | Unchanged | Holder Spring
 | Unchanged | Retaining Ring
New Previous Component

## Slide 12: HR2 Exhaust Assembly
1” Stiffening Doubler
1” Flange

## Slide 13: Standardized Construction
Exhaust Stack Mounting
  #10-32 x ¾” Bolts / Washers / nuts (28)
  Stiffening doubler and Gasket
  Insulation
  0.036 +/- 0.003 Aluminum Covering
  No substitutions permitted
Main body
  Dimensions
  Observation window
  Inner radiation doors
  Rear globar pan
  Reflector plate
  Diamond plate shaft
  Upper Pilot/Calibration Burner
  Second Stage Plate
    Hole pattern (Centered in Chamber)
    Perimeter sealed
    Removable

## Slide 14: Standardized Construction
Holding Chamber
  Dimensions
  Outer door
  Preheat sample position
Specimen Holder / Components
  Dimensions
  Retaining Ring
  Spring Plate
  Retaining Rod
  Drip Pan
  Millboard (Calibration only)
HFG Mounting
  Refractory Board (Thickness, HFG flush mount)
  Corner HFG position

## Slide 15: Standardized Thermopile / Globars
Thermopile mV Generating System
  New Thermocouples (9)
  Detailed Wiring Diagram
  72 inch Overall Maximum Length Limit
  24 Gauge, Solid Conductor, Thermocouple Grade Wire
  Quick Disconnects
  Cold Junctions Centrally Located in Lower Plenum: 1.5 inch Spacing
  Reference Junction: Closest To Air Inlet
Globars
  Construction: CHZ, Length, Diameter, Resistance
  Rheostat control of upper/lower pairs or individually
  Advisory on final power adjustments / Uniformity
  Voltage Fluctuation Tolerance (Globars)

## Slide 16: New: Upper pilot HSI (Hot Surface Igniter)
A 0.125 ± 0.005 inch diameter ceramic rod 8.0 ± 0.0625 inches in length is positioned directly in the flames of the upper pilot burner.
The rod is continuously heated by the flamelets acting as a hot surface igniter to auto-ignite any upper pilot flames should they go out.
The distance from the centerline of the burner tube to the centerline of the HSI rod is 0.75 ± 0.125 inch.

## Slide 17: Upper pilot HSI (Hot surface igniter)
HSI / Upper Pilot Burner Operation

## Slide 18: New Procedures: Heat Flux
HFG Type: Replaced Gardon with Schmidt-Boelter
HFG Guidance: Paint / Care / Mounting
New insitu method (Center)
  3.65 ± 0.05 W/cm2
New Uniformity Tolerance (4-Corners)
  3.65 +/- 0.10 W/cm2

## Slide 19: New Procedures: Heat Flux
Stability criteria
The term “stable” as it relates to heat flux density and chamber equilibrium is calculated using a moving average and expressed in terms of % standard deviation over a defined time period. Typically, stable conditions can be achieved between 60 and 90 minutes after heating has begun.
Heat flux (60 second moving average)
Heat flux gauge millivolt signal that varies less than 2.0% standard deviation over the last 60 seconds and having a calculated heat flux density that is within range.
Chamber equilibrium (15 minute moving average)
Thermopile millivolt signal that varies less than 2.0% standard deviation over the last 15 minutes commencing no sooner than 30 minutes after turning the heating elements on.

## Slide 20: Calibration Factor / HRR Equation
OSU – FAA Fire Test Handbook
Simplifying this gives…
Heat Release Rate = kW/m2
HR2 – Chapter HR
Simplifying this gives…
Heat Release Rate = kW/m2
Deleted
Sample Area Moved to HRR Calculation

## Slide 21: HR2: Total Heat Release Calculation
Total Heat Release = kW · min/m²
Simplifying this gives…
Total Heat Release kW · min/m²
Where:
n = Total number of data collection points in 2 minutes
x = Total number of data collection points in 1 minute
Average thermopile millivolt reading at baseline (mV)
Thermopile millivolt reading at any point during test (mV)
Calibration factor
Note: This formula uses the Riemann Left Hand Sum method of integration.

## Slide 22
Randomize 4 main parameters
DOE Test Plan Concept
Parameter | DESCRIPTION | Min. | Avg. | Max.
System Air Flow rates | SCFM | 19 | 20 | 21
Heat Flux (W/cm2) | Center | 3.60 | 3.65 | 3.70
Upper Pilot | Air (SLPM) | 0.98 | 1.0 | 1.2
 | Methane (SLPM) | 1.3 | 1.5 | 1.7

## Slide 23: Changes as result of DOE Testing
  Chamber Airflow:
  From: 20.0 ± 5% (19.0 – 21.0 SCFM)
  To: 20.0 ± 2% (19.6 – 20.4 SCFM)
  Upper Pilot Methane Flow:
  From: 1.50 ± 13.3% (1.30 – 1.70 SLPM)
  To: 1.50 ± 2% (1.47 – 1.53 SLPM)
  Upper Pilot Air Flow:
  From: 1.0 ± 20% (0.80 – 1.20 SLPM)
  To: 1.0 ± 5% (0.95 – 1.05 SLPM)
  Required inlet air be controlled by a Mass Flow Controller
  Replaced the panel mounted air flow meter (for upper pilot mixing air) with a MFC
  Potential Changes to Calibration Method (From Step to Ramp)

## Slide 24: Calibration
  Goals:
  Achieve better repeatability between successive calibrations
  Achieve good reproducibility between machines
  Reduce time req’d to complete the calibration process
  Reduced use of Methane Gas
  Make it easier

## Slide 25: Calibration
Replaced Wet Test Meter with Mass Flow Controller (MFC)
Upper Pilot Burner Methane MFC
  Thermal-based
  0.25 inch inlet/outlet fittings
  Calibrated annually with NIST traceability
  Calibrated for Methane gas
  Accuracy: 1% Full Scale
  STP 0ºC at 760 mmHg
  Signal Output To Data Acquisition System
Dual Purpose Burner (Upper)
  Calibration Burner (Replaces Calibration T-Burner)
  Upper Pilot Burner during testing - 1.5 SLPM Methane / 1.0 SLPM Air
  Note: Mixing air MFC has same specifications (calibrated for air)

## Slide 26: Calibration
Calibration Method Changes
  Optimized flow settings from maximum of 8 L/min down to 4 L/min
  Entire calibration sequence defined (Beginning to end)
  Exhaust ventilation system must be on
  Sample holder, with drip pan and millboard, must be in place
  Preheat at 4 SLPM for 3 minutes
  Step change in flow (1-2, 1-3 & 1-4 SLPM) for each 3 minute period
  Average final 10 seconds of data
  The average calibration factor is calculated
Calibration Pass / Fail Criteria
  Removed the 5% STDEV error between calibration constants (5)
  The calibration factor is determined using:
  Added Range: Calibration factor must be 0.085 ± 0.01 kW/mV

## Slide 27: New Procedures: Calibration
Calibration Method Changes
  Exhaust ventilation system must be on
  Sample holder, with drip pan and millboard, must be in place
  New 20 second moving average of the thermopile signal
  Preheat at 4 SLPM until a mV Threshold Limit is achieved
  Ramp down gas flow from 4-1 SLPM over a certain period
  Methane flow and real-time Thermopile signal are recorded
  A slope is generated (L/mV) to calculate the calibration factor
Calibration Pass / Fail Criteria
  The slope is used to determine the calibration factor using:
  Calibration factor must be 0.085 ± 0.01 kW/mV

## Slide 28: Chapter HF (Placeholder Document)
Scope
Definitions
Calibration Factor () Heat Release Rate
Total Heat Release Heat Flux Density
Thermopile Mass Flow Controllers
Methane Gas (CH4) NIST
NIST Traceability Stability
Standard Temp. and Press. (STP)
Principle of the Method

## Slide 29: Chapter HF (Placeholder Document)
Apparatus
Heat Release Rate Apparatus
Holding Chamber
Environmental Chamber
Pyramidal Section
Apparatus Components
Thermopile
Pilot Burners
Hot Surface Igniter (HSI)
Electronic Igniter System
Air Distribution System
Specimen Holder
Calibration Equipment

## Slide 30: Chapter HF (Placeholder Document)
Data Acquisition System/Display Devices
Test Specimen
Specimen Number
Specimen Size
Specimen Preparation
Specimen Orientation
Conditioning
Calibration Procedure of Equipment
Heat Flux Calibration (Center)
Heat Flux Uniformity Calibration (Corners)
Determination of the Calibration Factor ()

## Slide 31: Chapter HF (Placeholder Document)
Preparation/Performance of Test
Preparation
Performance of Test
Preparation of Further Test Runs
Presentation of Results
Heat Release Rate
Total Heat Release
Accuracy, Units and Rounding
Requirements
Peak Heat Release Rate (65 kW/m2)
2-Minute Total Heat Release (65 kW · min/m²)
Quantity of Samples (80% must pass)
Test Report

## Slide 32: Chapter HF: Calibration Requirements and Frequency
Equipment
  Heat Flux gauges
  Data Acquisition System/Display Devices
  Main Air Supply MFC
  Methane / Mixing Air MFC
Procedures
  System Heat Flux
  Heat Flux Uniformity
  Thermopile (Gas Calibration)

## Slide 33: Chapter HF (Update)
- Added “Burr-free holes , internally and externally, but not chamfered” requirement to:
  Upper pilot burner holes (15)
  Lower Plate holes (8)
  Second Stage Plate holes (120)
Calibration
  Valid Calibration Factor range: 0.085 +/- 0.010 kW/mV
  Drip pan / Millboard inserted during calibration
  Removed references to:
  Calibration Pass/Fail Criteria (5% STDEV)
  Calibration Factor Term: (kW/m2)/mv
- Standardized use of sample holder drip pans during testing.
  Required for all testing regardless of whether a material drips or not
  Optional: Permanently mounted to holder or removable

## Slide 34: Chapter HF Appendix
  Contains information and supplemental material pertinent to referenced paragraphs.
Tolerance table
System Overview Drawing (assembled)
Individual Components - 30 Detailed Drawings
  Methane Gas (CH4) Purity
  Heat Release Rate Apparatus; Insulation / Gasket Material
  Radiation Source
  Thermopile / Thermocouples
  Lower & Upper Pilot/Calibration Burner and Hot Surface Igniter (HSI)
  Igniter System

## Slide 35: Chapter HF Appendix
  Contains information and supplemental material pertinent to referenced paragraphs.
  Air Distribution System; Main Distribution Plate / Second Stage Plate
  Specimen Holder, Drip Pan Holding Chamber/Injection Mechanism
  Heat Flux Sensors
  MFC’s
  Specimen Preparation – Foil wrapping, 5 second rule and general cleaning procedures
  Heat Release Rate / Total Heat Release Calculations
  Nominal Operating Parameter Table

## Slide 36: HR2 Prototype Testing
Materials (5 coupons each)
Light Brown Panel (2012 RR)
Schneller Panel (2012 RR)

## Slide 37
Chart: Light Brown Honeycomb Panel Peak HRR vs. % STDEV Avg = 54 kW/m2; 16% STDEV
| A06 | B17 | A20 | B10 | B12 | B14 | A22 | A09 | A17 | A21 | A24 | A25 | A16 | B08 | A03 | B11 | B07 | B15 | B09 | A18 | A26 | A10 | A01 | B03 | A05 | B05 | A11 | A04 | A07 | A23 | B04 | A13 | A14 | A02 | A12 | B06
 | 40.3333 | 42.6667 | 43.2667 | 43.93 | 44.96 | 45.54 | 45.6333 | 46.3333 | 46.8433 | 48.37 | 49.4 | 49.4 | 49.4877 | 49.8 | 49.8 | 50.9867 | 52.02 | 52.7133 | 53.1933 | 53.8333 | 54.4967 | 54.7104 | 55.4 | 55.9 | 58.5067 | 59.5167 | 59.5533 | 61.17 | 63.476 | 63.9667 | 65.3 | 66.0767 | 66.1 | 67.3801 | 67.6967 | 72.6133
 | 0.02746 | 0.027063 | 0.021845 | 0.044468 | 0.043006 | 0.10153 | 0.025721 | 0.024922 | 0.0084488 | 0.0069157 | 0.0057256 | 0.17127 | 0.076911 | 0.013168 | 0.025044 | 0.035093 | 0.05967 | 0.051233 | 0.023226 | 0.14344 | 0.077249 | 0.048396 | 0.067672 | 0.033492 | 0.06748 | 0.034494 | 0.089224 | 0.053982 | 0.066734 | 0.02208 | 0.015354 | 0.010311 | 0.10043 | 0.073553
HR2: 53.5

## Slide 38
Chart: Light Brown Honeycomb Panel Total HR vs. % STDEV Avg = 72 kW*min/m2; 13% STDEV
| A06 | A10 | B07 | A09 | B14 | A20 | B11 | B10 | B12 | B17 | B05 | A17 | A22 | A23 | A16 | B15 | B09 | A02 | A25 | A24 | A01 | A21 | A03 | B08 | A14 | A05 | A04 | B03 | A18 | A26 | B04 | A11 | A07 | A12 | B06 | A13
 | 52.63 | 54.2936 | 58.7867 | 61 | 61.59 | 63.6667 | 63.9567 | 66.4767 | 66.7567 | 67 | 67.1633 | 67.4367 | 67.4533 | 67.6333 | 68.4447 | 68.6233 | 69.3867 | 69.8669 | 69.9 | 70.1 | 70.8 | 71.3 | 72.9 | 74.8 | 75.4333 | 76.8233 | 79.5533 | 79.58 | 81.0333 | 81.2467 | 81.5333 | 81.61 | 81.7207 | 82.9997 | 89.3867 | 94.97
 | 0.006593 | 0.076948 | 0.082661 | 0 | 0.044444 | 0.064423 | 0.085444 | 0.035564 | 0.069668 | 0.025852 | 0.042361 | 0.014271 | 0.066833 | 0.029878 | 0.08268 | 0.02287 | 0.039863 | 0.07379 | 0.032371 | 0.023519 | 0.01351 | 0.069673 | 0.011882 | 0.055545 | 0.049865 | 0.039664 | 0.024381 | 0.061972 | 0.01837 | 0.027144 | 0.072457 | 0.053465 | 0.0088049 | 0.014979
HR2: 69.6

## Slide 39
Chart: Dark Brown Honeycomb Panel Peak HRR vs. % STDEV Avg = 49 kW/m2; 12% STDEV
| A10 | B14 | B11 | A12 | A20 | A06 | B03 | A21 | A03 | A16 | A01 | A25 | B08 | B17 | A26 | A04 | A24 | B12 | B07 | A22 | A14 | B10 | A02 | A17 | A09 | B05 | B04 | A18 | A23 | A13 | B15 | B09 | B06 | A11 | A07 | A05
 | 36.1905 | 38.79 | 40.4 | 41.7333 | 42.2333 | 43.4 | 44.24 | 45.1967 | 45.4 | 45.499 | 45.9 | 46.1667 | 46.3333 | 46.3333 | 46.6567 | 47.4367 | 48.8 | 49.22 | 49.5467 | 49.5667 | 49.5667 | 50.6667 | 50.7799 | 50.9967 | 51.3333 | 51.53 | 51.8 | 52.1 | 52.3667 | 53.4667 | 53.74 | 54.0567 | 57.8767 | 58.6 | 58.7107 | 65.1733
 | 0.031082 | 0.027177 | 0.066728 | 0.14016 | 0.0068352 | 0.015815 | 0.17339 | 0.04127 | 0.015419 | 0.088492 | 0.024346 | 0.0075796 | 0.044928 | 0.091946 | 0.013006 | 0.089569 | 0.14295 | 0.021595 | 0.024238 | 0.024507 | 0.065051 | 0.050697 | 0.022494 | 0.029914 | 0.037234 | 0.0083664 | 0.059362 | 0.073839 | 0.062692 | 0.028966 | 0.048027 | 0.025613 | 0.024375 | 0.044468
HR2: 47.9

## Slide 40
Chart: Dark Brown Honeycomb Panel Total HR vs. % STDEV Avg = 36 kW*min/m2; 16% STDEV
| A10 | A16 | B05 | A06 | A01 | B11 | B14 | A02 | B03 | B07 | A12 | A23 | A17 | A24 | A25 | A21 | B12 | A20 | B15 | A22 | A09 | A04 | A14 | A03 | B17 | A18 | A26 | B04 | B10 | A07 | A11 | B09 | B08 | B06 | A13 | A05
 | 19.6117 | 26.163 | 27.565 | 27.6367 | 27.7 | 28.1967 | 32.5433 | 32.5763 | 33.15 | 33.25 | 33.7533 | 34.1333 | 34.1767 | 34.5 | 34.9 | 35.15 | 35.3767 | 35.6 | 36.3367 | 36.66 | 36.6667 | 36.7633 | 36.8667 | 36.9333 | 37 | 39.0733 | 39.4867 | 39.7333 | 41.07 | 41.2304 | 41.2967 | 41.73 | 41.9 | 46.06 | 46.1833 | 47.2833
 | 0.20075 | 0.16215 | 0.085935 | 0.14819 | 0.18534 | 0.065565 | 0.051466 | 0.27373 | 0.18211 | 0.085316 | 0.12908 | 0.056088 | 0.12663 | 0.095306 | 0.033231 | 0.035197 | 0.13875 | 0.066027 | 0.08767 | 0.060161 | 0.021925 | 0.033637 | 0.054054 | 0.14374 | 0.10825 | 0.016375 | 0.0058488 | 0.067242 | 0.071473 | 0.0301 | 0.03307 | 0.052547 | 0.13147 | 0.092172
HR2: 36.9

## Slide 41: Next
Continue work on New Calibration approach (Ramp method)
Additional Unit Buildups for Comparative Testing
Testing of other type materials?
  (Task group input needed)
Continue DOE as needed