## Slide 1: RECENT DEVELOPMENTS IN MICROSCALE COMBUSTION CALORIMETRY
Richard Walters, Richard Lyon, Louise Speitel, FAA
Natallia Safronava, TAMI
Haiqing Guo, CFar Services

## Slide 2: Microscale Combustion Calorimeter
Developed at FAA Technical Center
Measures Material Properties
Sample Size 5-20 mg
Temperature Resolved Data
Repeatable & Reproducible Results
Burning Rate
Area = Heat of Combustion
Ignition Temperature
PMMA

## Slide 3: Heat Release Rate Calculation

## Slide 4: Heat Release Rate Calculation
Case 1: ASTM Calculation Case 7: New Calculation
- - - - Literature Value

## Slide 5: Baseline Correction Method
=
= -
= -
-
Coefficients for Thermal Baseline Drift
Baseline Correction
Baseline Corrected Heat Release Rate

## Slide 6: Baseline Correction Method

## Slide 7: High Precision O2 Measurement
Paramagnetic Oxygen Sensor
  Less noise
  Eliminates baseline drift
  Longer lifespan
  More expensive

## Slide 8: Baseline Correction Method
Standard (chemical) oxygen sensor
ASTM D7309-13 calculation
Upgraded (paramagnetic) oxygen sensor
ASTM D7309-18 calculation
Current Practice
High Precision Method

## Slide 9: Fire Growth Capacity (FGC)
Measure Heat Release Rate Q vs Temperature T as per ASTM D7309 (5 replicates)
Integrate Q/ versus T to obtain Q vs T, i.e., Q(T )
Obtain Total Heat Release
Q(T∞) = Q∞ = hc(J/g)
Obtain T1 at 5% deflection from Q(T) baseline, i.e., at 0.05Q∞
Obtain T2 at Q∞-(0.05Q∞), i.e., 0.95Q∞.
Calculate Fire Growth Capacity (FGC)
0
10
20
30
40
50
60
70
80
0
2
4
6
8
10
200
300
400
500
600
700
800
900
Specific Heat Release Rate, Q (W/g)
Temperature, T (°C)
Q (J/g)
Q∞
T1
T2
-5% Q∞
+5%
Q∞
T0 = Standard Temperature = 25°C
T1  Ignition temperature
T2 = Burning temperature

## Slide 10: FGC

## Slide 11: Revisions To ASTM D7309 Standard
New HRR Calculation
New Baseline Correction Method
Interlaboratory Study
New Precision and Bias Statement
Fire Growth Capacity

## Slide 12: Interlaboratory Study
Needed for ASTM
New Precision & Bias Statement
Need Samples
  Aircraft Materials
  Composite (homogeneous & heterogeneous)
Need Participants

## Slide 13: Proposal for Component Substitution
ASTM D7309 (MCC)
Test and Criterion
Components
(mg)
Similar
MCC
sufficient
Different
FAR Test
of coupon
Similar
Different
Coupon
sufficient
FAR Test
of construction
(fabricated part)
Coupons (g)
(self supporting components)

## Slide 14: Incomplete Combustion Experiments
Extract more meaningful data from MCC
Measure effect of gas phase FR additives
Measure toxic gas production
103 kg
1 kg
10-6 kg

## Slide 15: Combustion Chemistry
Fuel + Some Oxygen  CO + HCN + Hydrocarbons + Other Toxics
Fires generate toxic combustion products when oxygen becomes depleted (vitiated),
Incomplete & toxic combustion products
Complete (less-toxic) combustion products
Excess Oxygen
Insufficient Oxygen
CO2 + H2O + N2
Late Stage Fire
Early Stage Fire

## Slide 16: High Temperature Combustor
Installed on MCC
  High temperature
  Small volume
Approaching flame temps
  1500°C tested so far
  Soot formed
Controlled Fuel:Oxygen ratios
  Incomplete combustion

## Slide 17: Controlled Fuel:Oxygen Ratio Experiments
 < 1 = Fuel Lean (excess oxygen)
 > 1 = Fuel Rich (oxygen starved)
F
=
mFuel / mO2
 is the fuel/oxygen mass ratio relative to the stoichiometric mass ratio.
(mFuel/mO2 )stoichiometric

## Slide 18: Controlling F in MCC
0
0.5
1
1.5
2
2.5
3
200
300
400
500
600
700
Percent of Oxygen Consumed, ∆ XO2
Sample Temperature, T (C)
Training Set
Program ( = 1)
Measured ( = 1)
Run material of mass m0 at  = 1 K/s, XO2 = 0.2, and Tc = 900C.
Measure ∆XO2/m0 at each sample temperature T for complete combustion of gases to obtain training set, (∆XO2/m,T)0.
Store training set (XO2,T,m)0 for material.
Use (∆XO2/m,T)0 data to control O2/N2 flows such that:
(∆XO2(T)/m)0
∆XO2(T)/m

=
=
m0
m
∆XO2(T)
(∆XO2(T))0

## Slide 19: Hyphenated MCC Techniques Using FTIR & Gas Analyzers

## Slide 20: Multi-Point Calibration Spectral Regions
CO
50 ppm
500 ppm
1000 ppm
3000 ppm
4000 ppm
5000 ppm
10000 ppm
20000 ppm
30000 ppm
40600 ppm

## Slide 21: FTIR Method

## Slide 22: Testo Portable Gas Analyzer
No HCN?

## Slide 23: Nylon Gas Analysis

## Slide 24: Comparison of All Vitiated PA66 Yields
0.01
0.1
1
10
100
0
0.5
1
1.5
2
2.5
HCN Yield (g/kg)
Equivalence Ratio
Hydrogen Cyanide
1
10
100
1000
0
0.5
1
1.5
2
2.5
CO Yield (g/kg)
Equivalence Ratio
Carbon Monoxide
Fire Propagation
Apparatus
Purser Furnace
1/3 ISO Room
ISO Room
Controlled Atm.
Cone Calorimeter
MCC-FTIR
MCC-CO analyzer
(A. Stec & R. Hull, Fire Toxicity, 2010)

## Slide 25: Fuel:Oxygen in Diffusion Flame

0
∞
Air
Fuel
Combustible Solid
Fuel Gases
Air
Laminar
Diffusion Flame
Varies from ∞ to 0 in diffusion flame
Distance into Flame Sheet
Fflame = Average Value
Soot
Formation

## Slide 26: Soot Microscopy - TEM
UMass Amherst
Polymer Science & Engineering
The Silvio O Conte National Center For Polymer Research

## Slide 27: Gravimetric Soot Experiments - PS
1000°C
1250°C
1375°C
1500°C
1
1.25
1.5
0.75
Equivalence Ratio, F
Filter Assembly
Stainless Steel Filter Housing
25mm PTFE unlaminated
0.2 micron membrane filters
Vacuum dried and weighed
prior to and after testing

## Slide 28: Gravimetric Soot Yields

## Slide 29: Summary & Conclusions
Improved accuracy and precision of MCC data
  HRR calculation
  Baseline
  O2 analyzer
Fire growth capacity
Revisions to ASTM standard
  ILS
Similarity determinations for small changes
Recreates flaming combustion products
  Toxicity measurements
  Gravimetric soot measurements

## Slide 30: Contact Info
Richard Walters
Federal Aviation Administration
Fire Safety Branch ANG-E21
Atlantic City Int’l Airport, NJ 08405 USA
richard.walters@faa.gov
609-485-4328