## Slide 1: Flow Visualization in the OSU
Triennial Conference on Cabin and Fire Safety
Tina Emami
October 28-31, 2019

## Slide 2: Airflow Project Goals
Create an adjustment to the current OSU apparatus in order to create an even and reproducible air flow.

## Slide 3: Airflow Project Background
Two areas are being examined
Area 1: Lower plenum
  Airflow is measured
  Changes are made in this area
Area 2: Sample area
  Airflow is measured here to understand the effect of the changes made to the lower plenum area

## Slide 4: Original Test Setup
CAMERA
LASER
OSU is oriented so the glow bars are facing the camera
Chamber 1

## Slide 5: Area 1 (Lower Plenum) Location Original Test Setup
The green squares in the lower plenum are the Area 1 location measurement

## Slide 6: Area 1 (Lower Plenum) Original Test Setup
Results shown are averages of 500 data images

## Slide 7: Area 2 Location (Near Sample)

## Slide 8
Measurement Location

## Slide 9: Area 2 (Near Sample) Original OSU Setup

## Slide 10: New Setup 1 Comparison Measured at Test Sample
New Set Up 1
New Test Setup 1
Original OSU Setup

## Slide 11: New Setup 2 Comparison Measured at Test Sample
New Test Setup 2
Original OSU Setup

## Slide 12: Reticulated Foam when inside the hot OSU

## Slide 13: Other Concepts Studied

## Slide 14: New Setup 5 Comparison Air Distribution Tube Measured at Test Sample
Original OSU Setup
New Test Setup 5

## Slide 15: New Setup 5

## Slide 16: New Tube Placed in OSU

## Slide 17: Analyzing the New Air Distribution Tube
Airflow measurements using a vane anemometer and hot wire anemometer
Temperature readings before and during sample testing
Heat release test comparison before and after the addition of the tube

## Slide 18: Preliminary Measurements
Measurements taken without any changes in a metal OSU
While the unit is cold, air velocity was measured in the OSU with a vane anemometer
Left:
0.44 m/s
Right:
0 m/s
Air
Inlet
Center:
0.32-0.40 m/s

## Slide 19: Airflow measurements with a hot wire anemometer
Simple airflow measurements were performed in a cold, metal OSU
All doors are open during measurement, because the hot wire anemometer will not allow the doors to close fully

## Slide 20: Airflow measurements with a hot wire anemometer Air Distribution Tube inside OSU
Glow Bars
Viewing Window
Doors
0.05 m/s
0.11 m/s
0.40 m/s
0.13 m/s
0.20 m/s
0.28 m/s
0.25 m/s
0.18 m/s
0.25 m/s
Key
0.05 m/s
0.11 m/s
0.25 m/s
0.18 m/s
0.4 m/s

## Slide 21: Baseline Tests for Temperature
Each red dot indicates a temperature measurement taken.
Each blue dot is a side view of an existing red dot, to clarify its location in space.

## Slide 22: Rear Inner Cone Temperatures – Left and Right Baseline Measurement
Before adding the air distribution tube, the left and right corners of the Rear Inner Cone had an average temperature difference of 35°F.
After the air distribution tube, they have an average temperature difference of 0.35°F.
Addition of tube increased the Rear Cone temperature by 149°F on the left and 115°F on the right.

## Slide 23: Forward Inner Cone Temperatures – Left and Right Baseline Measurement
Before adding the air distribution tube, the left and right corners of the Forward Inner Cone had an average temperature difference of 34°F.
After the air distribution tube, they have an average temperature difference of 1°F.
Addition of tube increased the Inner Rear Cone temperature by 159°F on the left and 126°F on the right.

## Slide 24: Inner Cone Temperatures – Left and Right Baseline Measurement
Before the air distribution tube, the four corners of the inner cone were all averaged at 456°F with a standard deviation of 23.1%
The center of the cone read an average of 422°F
After the addition of the air distribution tube, the four corners of the inner cone are all averaged at 593°F with a standard deviation of 5.24%
The center of the cone read an average of 588°F
The result shows a more stable and uniform temperature reading,
but at a significant increase in temperature.

## Slide 25: Mask Temperature Baseline Measurement
After the addition of the air distribution tube, the average center mask temperature decreased by 62°F.
Chart: Mask Center Temperature - Pre vs Post Air Tube

## Slide 26: Baseline Tests in the OSU
Four different honeycomb samples were tested in the OSU for comparative purposes
Each sample was tested three times
Sample | Thickness
Sample A | 1/8 in
Sample B | 11/16 in, white decorative
Sample C | 7/8 in, no decorative
Sample D | 1/4 in, white decorative

## Slide 27: Mask Temperature Samples A and B
Chart: Sample B - Center Mask Air Temperature
Chart: Sample A - Center Mask Air Temperature

## Slide 28: Exhaust Air Temperature – Samples A and B
Chart: Sample A - Exhaust Air Temperature
Chart: Sample B - Exhaust Air Temperature
Sample A and B were tested before and after the air distribution tube was placed.
After the addition of the air distribution tube, the exhaust air temperature increased by 24°F for sample A and 23°F for sample B

## Slide 29: Testing Sample A – Heat Release
Chart: Sample A - HRR Post Air Dist. Tube
Chart: Sample A - HRR Pre Air Dist. Tube
Chart: Sample A - HRR Pre vs Post Air Dist. Tube
 | Pre Air Dist. Tube | Post Air Dist. Tube
2 min. Total Heat Release kW min/m^2 | 33.30 | 33.27
Peak Heat Release kW/m^2 | 44.93 | 48.11

## Slide 30: Testing Sample B – Heat Release
Chart: Sample B - HRR Post Air Dist. Tube
Chart: Sample B - HRR Pre Air Dist. Tube
Chart: Sample B - Pre vs Post Air Dist. Tube
 | Pre Air Dist. Tube | Post Air Dist. Tube
2 min. Total Heat Release kW min/m^2 | 37.18 | 44.64
Peak Heat Release kW/m^2 | 39.01 | 39.60

## Slide 31: Summary of Results
The cone now has an overall stable and uniform temperature reading throughout but at the cost of a significantly hotter temperature.
The exhaust temperatures increased while the temperature in the mask area decreased.
  These bring us to the educated guess that the air velocity near the sample holder actually did decrease.
The 2 minute Total Heat Release increased for every sample except sample A which stayed the same.
The Peak Heat Release varied throughout the tests.
Samples B, C, and D had worse variability in their Heat Release than they did before the Air Distribution Tube.

## Slide 32: Problems encountered
Plexiglas could be affecting these readings through static electricity
The velocity readings were shown much slower than what was measured with anemometers in the traditional OSU
Since the PIV was displaying the average velocity, and there was a lot of circulation, it could have been measuring upwards flow half the time and downwards flow half the time, averaging to very slow final speeds.

## Slide 33: Future Work
Take out the air distribution tube and re-measure the original OSU with the hot wire anemometer to see if there are any differences from the first time
Will continue this work with hot wire anemometers and larger scale flow measurements and analysis

## Slide 34: Questions or Comments?
Tina Emami
FAA William J. Hughes Technical Center
ANG-E212 Materials Fire Test
Building 275
Atlantic City Int’l Airport, NJ 08405
tina.emami@faa.gov
609-485-4277