## Slide 1: HEAT RELEASE RATE Updates 2014 June Materials Meeting Switzerland
Materials Working Group
Michael Burns, FAA Tech Center
June, 2014

## Slide 2: AGENDA
OSU
  Orifice Meter Discussion
  (Static Pressure Taps)
Chapter HR Update (hr2)
  Combined Supplement with Test method

## Slide 3: OSU:
  Protruding pressure taps (beyond the pipe inner wall) vs. surface pressure taps (flush with inner wall) may have a influence on pressure readings
  Could have an influence on total airflow

## Slide 4: HR2: HEAT RELEASE TEST METHOD
  Supplemental material moved to Test Method as follows:
  Added BTU/ft2*sec units
  Environmental Chamber – Added Radiation Door Assembly Figure
  Radiation Source
    Power Supply Capacity – 12 kVA is sufficient
    Diamond-Shaped Mask Mounting – ½ inch diameter machine screw
    (approximately 4 inches in length)
    Mask Position – Approximately 3 inches from the reflector plate
    Mask shaft-head or nut only on side facing specimen
    (no threads extending towards specimen)

## Slide 5: HR2: HEAT RELEASE TEST METHOD
  Thermopile Wiring
    Color coding differences - Chromel yellow/green; Alumel red/white
  Upper Pilot Burner
    Caution Statement – Do not disturb burner position once set
    Flame Profile – To increase orange tip, reduce air from the mixture
  Drill hole sizes added for:
    Upper Pilot Burner (#59 Drill)
    Air Distribution Plate (# 4 Drill)
    Second Stage Plate (# 28 Drill)
  Drip Pan - Foil can be used to line the drip pan to facilitate cleaning after use

## Slide 6: HR2: HEAT RELEASE TEST METHOD
  Heat Flux Gauge (HFG)
    Cooling Water Tubing – Rigid due to the high temperatures
    Cleaning – Gently wipe with soft, water-dampened sponge
    Caution Statement – Damage done to the coating during the cleaning process will affect the measurement accuracy of the sensor
  Specimen Preparation
    Added recommended specimen foil wrapping technique
    Foil trimming – no scoring or incisions permitted
    Spring Plate Pressure – Recommended applying equal pressure to specimens of equal thickness by inserting the specimen retaining rod in the same holes for each specimen tested

## Slide 7: HR2: HEAT RELEASE TEST METHOD
  Heat Flux Calibration Procedure (Corner)
    Caution Statement – Do not disturb burner position when rotating calibration fixture 90 degrees
    Added options for improving 4-corner uniformity:
      Adjust power settings
      Reposition globars - Hottest to coolest (bottom to top)
      Move mask – Closer to or further from HFG
      Move globars – Adjustment left or right
      Reflector plate – Adjust slope (top and bottom)
  Added Rounding, Accuracy and Units section

## Slide 8: HR2: HEAT RELEASE TEST METHOD
  Calculation of Calibration Factor Kh – Added example assuming 60 mV baseline; 7 mV rise for each change in flow

## Slide 9: HR2: HEAT RELEASE TEST METHOD
  Test Preparation
    Thermopile – Baseline mV observed daily as an indicator of developing problems
      Variations in heat flux (heat flux gauge malfunction)
      Defects in thermopile
      Air leakage – Environmental chamber/lower plenum, exhaust stack or holding chamber door seal
  Test Performance
    The use of an externally positioned mirror may assist in viewing upper pilot flames during testing

## Slide 10: HR2: HEAT RELEASE TEST METHOD
  Heat Release Rate – Added example for calculating HRR (Peak HRR)

## Slide 11: HR2: HEAT RELEASE TEST METHOD
  Total Heat Release – Added example for calculating Total HR (2-Min. THR)

## Slide 12: HR2: HEAT RELEASE TEST METHOD
  Test Report
    Laboratory
    Fully identify the material tested
    Graphically report the heat release rate (kW/m2) as a function of time (in seconds) for each test
    Maximum heat release rate and time (in seconds) it occurs during each test
    Total heat released during the first 2 minutes of each test
    Radiant heat flux to the specimen

## Slide 13: HR2: HEAT RELEASE TEST METHOD
  Test Report (Continued)
    Calibration factor Kh (kW/mV )
    Average mV baseline for each specimen tested
    Melting, sagging, delaminating, or other behavior that affected the exposed surface area or mode of burning that occurred and the time(s) at which such behavior occurred
    Total number of materials tested
    Total number of materials passed

## Slide 14: HR2: HEAT RELEASE TEST METHOD
  Requirements
    Quantity of Samples – Minimum of three samples must be tested and 80% or greater must pass

## Slide 15: QUESTIONS?
  What is full of holes, but can still hold a lot of water?
  Answer: Sponge