## Slide 1: HEAT FLUX GAUGE CALIBRATION UPDATE (CHAPTER HF DRAFT) 2014 June Materials Meeting Switzerland
Materials Working Group
Michael Burns, FAA Tech Center
June, 2014

## Slide 2: AGENDA
Calibration of Secondary Standard HFG
Coating
HFG Mounting
DAQ / Calibration Software
Calibration Procedure
Analysis
Reporting

## Slide 3: CHAPTER HF DRAFT – CALIBRATION
Secondary Standard HFG
A minimum of 4 data points are required encompassing nominal heat flux values of 0.88, 1.76, 2.64 and 3.52 BTU/ft2*sec (1, 2, 3 and 4 W/cm2).
The zero flux data point must be omitted. Zero offset data is not to be considered and data beyond these ranges must be disregarded.

## Slide 4: CHAPTER HF DRAFT – CALIBRATION
Coating
The HFG must have a thin, full-faced, opaque coating of high quality, high-temperature, ultra-flat black paint having a diffuse surface finish.
The nominal dry film thickness of the coating on the sensor face must be 2 mils or less with the “orangey” copper face of the sensor not visible.
Note: Mankiewicz Black Suede, Krylon Flat Black, Pyromark 1200 Flat Black paints or products of similar specification or function have been found suitable.

## Slide 5: CHAPTER HF DRAFT – CALIBRATION
Data Acquisition System Specification
The calibration software must convert millivolt data of a Working HFG (gauge to be calibrated) to that of a Standardized gauge (having a known slope) in order to generate a calibration factor.
Calibration data between the range of 0.88 BTU/ft2*sec (1 W/cm2) and 3.52 BTU/ft2*sec (4 W/cm2) is used to linearize the slope.
Zero offset data is not to be considered and data beyond these two ranges must be omitted.

## Slide 6: CHAPTER HF DRAFT – CALIBRATION
Mounting HFGs for Calibration
Each HFG must be mounted in a supporting device protruding through 0.5 inch (12.7 mm) thick rigid refractory board.
Each gauge has a 180º field of view of the radiant heat source therefore the surface must be mounted flush with the insulation board and not recessed.
Note: Kaowool M-Board, Superwool 607 or product of similar specification or function has been found suitable.

## Slide 7: CHAPTER HF DRAFT – CALIBRATION
Calibration Procedure
STEP 2
Apply radiant heat to both sensors until the Standardized HFG reading reaches 4.4 BTU/ft2*sec (5 W/cm2). A heat radiance increase from 0 to 4.4 BTU/ft2*sec (0 to 5 W/cm2) must occur within 45 ± 5 seconds to minimize heating of surrounding materials and/or equipment.
STEP 3
Immediately begin decreasing the radiant heat. A cool down rate from 4.4 BTU/ft2*sec (5 W/cm2) to less than 0.88 BTU/ft2*sec (1 W/cm2) must occur over a minimum time period of 90 seconds.

## Slide 8: CHAPTER HF DRAFT – CALIBRATION
Calibration Procedure (Continued)
STEP 4
At approximately 3.52 BTU/ft2*sec (4 W/cm2) begin recording signal data continuously for both HFGs.
STEP 5
At approximately 0.88 BTU/ft2*sec (1 W/cm2) stop recording signal data.

## Slide 9: CHAPTER HF DRAFT – CALIBRATION
Analysis
Multiply the Standardized HFG millivolt values by the known slope and calculate the heat flux level for each data point. A least squares fit procedure is performed using the calculated heat flux and the Working HFG millivolt values. The linearized slope is determined by:
∆Y (Units) / ∆X (mV)
Where:
Y = Standardized HFG (BTU’s/ft2*sec or W/cm2)
X = Working HFG (millivolts)
Note:
To convert BTU’s/ft2*sec to W/cm2: multiply slope by 1.135654
To convert W/cm2 to BTU’s/ft2*sec: multiply slope by 0.88

## Slide 10: CHAPTER HF DRAFT – CALIBRATION
Calibration Reporting Parameters

## Slide 11
TOPIC | SUBTOPIC | DETAILS | 
Facility conducting calibration |  |  | 
Customer | Lab performing future tests with the Working HFG |  | 
Calibration date | Year / Month / Day |  | 
Calibration expiration date | Year / Month / Day |  | 
Working HFG Information | Manufacturer |  | 
 | Part number |  | 
 | Serial number |  | 
 | Gauge type | Gardon | 
 |  | Schmidt-Boelter | 
 | Coating | Paint Information | 
 |  | Nominal Thickness (mils) | 
 | Calibration Factor | BTU/ft2*sec | 
 |  | W/cm2 | 
 | Voltage / Resistance (Zero Flux) | Pre Calibration | Voltage
 |  |  | Ohm
 |  | Post Calibration | Voltage
 |  |  | Ohm
 | Prior Calibration Data (If available) | Date | 
 |  | Calibration Factor | BTU/ft2*sec
 |  |  | W/cm2
 |  | Facility where calibration was conducted | 
 |  | Percent change in calibration factor | 
 |  | Traceability records to a Standardized HFG | 

## Slide 12
TOPIC | SUBTOPIC | DETAILS | 
Standardized HFG Information | Manufacturer |  | 
 | Part number |  | 
 | Serial number |  | 
 | Gauge type | Gardon | 
 |  | Schmidt-Boelter | 
 | Coating | Paint information | 
 |  | Nominal thickness (mils) | 
 | Calibration Factor | BTU/ft2*sec | 
 |  | W/cm2 | 
 | Voltage / Resistance (Zero Flux) | Pre Calibration | Voltage
 |  |  | Ohm
 |  | Post Calibration | Voltage
 |  |  | Ohm
 | Calibration Date of Standardized HFG | Year / Month / Day | 
 | Traceability Records to NIST Primary Standard HFG |  | 
HFG cooling water | Temperature | ̊F | 
 |  | ̊C | 
XY plot or raw data table |  |  | 
Linearity of HFG mV signals during calibration (Confidence level) |  |  | 
Calibration Traceability |  |  | 
Calibration Certificates of Devices | (DAQ, NIST etc. ) |  | 

## Slide 13
Questions?