## Slide 1: Flow Visualization in the OSU
IAMFTF, Cologne, Germany
Tina Emami
June 18, 2019

## Slide 2: Airflow Project Goals
Create an adjustment to the current OSU apparatus in order to create an even and reproducible air flow.

## Slide 3: Test Setup 1
CAMERA
LASER
OSU is oriented so the glow bars are facing the camera
Chamber 1
Front of OSU

## Slide 4: Measurement Locations- From Test Setup 1

## Slide 5: Bottom Chamber – Front View Original Test Setup
Results shown are averages of 500 data images

## Slide 6
Measurement Location

## Slide 7: Location of Current Laser Light Sheet

## Slide 8: New Test Setup
PVC tube with two rows of evenly spaced 5/8” holes 45° from horizontal was constructed
Cap is glued to the end of the tube
Tube is placed in the bottom chamber of the OSU
Holes will be oriented to face downwards in order to more evenly mix the flow

## Slide 9: New Test Setup

## Slide 10: 5/8- inch Holes in Air Distribution Tube
Air
Inlet
Here, we are measuring by the sample face again

## Slide 11: Baseline Tests in the OSU
Four different honeycomb samples were tested in the OSU for comparative purposes
Each sample was tested three times
Sample | Thickness
Sample A | 1/8 in
Sample B | 11/16 in, white decorative
Sample C | 7/8 in, no decorative
Sample D | 1/4 in, white decorative

## Slide 12: Baseline Tests
Each red dot indicates a temperature measurement taken.
Each blue dot is a side view of an existing red dot, to clarify its location in space.

## Slide 13: New Tube Placed in OSU

## Slide 14: Animation of New Setup

## Slide 15: Vane Anemometer Measurement Location (Cold OSU)

## Slide 16: Vane Anemometer Measurements (Cold OSU)
Before air distribution tube
After air distribution tube
0.42 m/s
0.32-0.4 m/s
fluctuation
0 m/s
0.53 m/s
0.18 m/s
0 m/s

## Slide 17: Pressure Measurements (Cold OSU)
Before air distribution tube
After air distribution tube
10.8 INWC
0.4 INWC
10.2 INWC
0.4 INWC

## Slide 18: Rear Inner Cone Temperatures – Left and Right Baseline Measurement
Before adding the air distribution tube, the left and right corners of the Rear Inner Cone had an average temperature difference of 35°F.
After the air distribution tube, they have an average temperature difference of 0.35°F.
Addition of tube increased the Rear Cone temperature by 149°F on the left and 115°F on the right.

## Slide 19: Forward Inner Cone Temperatures – Left and Right Baseline Measurement
Before adding the air distribution tube, the left and right corners of the Forward Inner Cone had an average temperature difference of 34°F.
After the air distribution tube, they have an average temperature difference of 1°F.
Addition of tube increased the Inner Rear Cone temperature by 159°F on the left and 126°F on the right.

## Slide 20: Inner Cone Temperatures – Left and Right Baseline Measurement
Before the air distribution tube, the four corners of the inner cone were all averaged at 456°F with a standard deviation of 23.1%
The center of the cone read an average of 422°F
After the addition of the air distribution tube, the four corners of the inner cone are all averaged at 593°F with a standard deviation of 5.24%
The center of the cone read an average of 588°F
The result shows a more stable and uniform temperature reading,
but at a significant increase in temperature.

## Slide 21: Mask Temperature Baseline Measurement
After the addition of the air distribution tube, the average center mask temperature decreased by 62°F.
Chart: Mask Center Temperature - Pre vs Post Air Tube

## Slide 22: Mask Temperature Samples A and B
Chart: Sample B - Center Mask Air Temperature
Chart: Sample A - Center Mask Air Temperature

## Slide 23: Exhaust Air Temperature Baseline Measurement
Baseline temperature measurements were taken of the OSU with the glow bars on and pilot burners lit.
After the addition of the air distribution tube, the average exhaust temperature increased by 56°F.
Chart: Exhaust Air Temperature Pre vs Post Air Tube

## Slide 24: Exhaust Air Temperature – Samples A and B
Chart: Sample A - Exhaust Air Temperature
Chart: Sample B - Exhaust Air Temperature
Sample A and B were tested before and after the air distribution tube was placed.
After the addition of the air distribution tube, the exhaust air temperature increased by 24°F for sample A and 23°F for sample B

## Slide 25: Exhaust Air Temperature – Samples C and D
Sample C and D were tested before and after the air distribution tube was placed.
After the addition of the air distribution tube, the exhaust air temperature increased by 32°F for sample C and 20°F for sample D
Chart: Sample C - Exhaust Air Temperature
Chart: Sample D - Exhaust Air Temperature

## Slide 26: Testing Sample A – Heat Release
Chart: Sample A - HRR Post Air Dist. Tube
Chart: Sample A - HRR Pre Air Dist. Tube
Chart: Sample A - HRR Pre vs Post Air Dist. Tube
 | Pre Air Dist. Tube | Post Air Dist. Tube
2 min. Total Heat Release kW min/m^2 | 33.30 | 33.27
Peak Heat Release kW/m^2 | 44.93 | 48.11

## Slide 27: Testing Sample B – Heat Release
Chart: Sample B - HRR Post Air Dist. Tube
Chart: Sample B - HRR Pre Air Dist. Tube
Chart: Sample B - Pre vs Post Air Dist. Tube
 | Pre Air Dist. Tube | Post Air Dist. Tube
2 min. Total Heat Release kW min/m^2 | 37.18 | 44.64
Peak Heat Release kW/m^2 | 39.01 | 39.60

## Slide 28: Testing Sample C – Heat Release
 | Pre Air Dist. Tube | Post Air Dist. Tube
2 min. Total Heat Release kW min/m^2 | 42.51 | 47.79
Peak Heat Release kW/m^2 | 43.56 | 58.95

## Slide 29: Testing Sample D – Heat Release
 | Pre Air Dist. Tube | Post Air Dist. Tube
2 min. Total Heat Release kW min/m^2 | 33.64 | 40.66
Peak Heat Release kW/m^2 | 44.99 | 37.09

## Slide 30: Summary of Results
The cone now has an overall stable and uniform temperature reading throughout but at the cost of a significantly hotter temperature.
The exhaust temperatures increased while the temperature in the mask area decreased.
  These bring us to the educated guess that the air velocity near the sample holder actually did decrease.
The 2 minute Total Heat Release increased for every sample except sample A which stayed the same.
The Peak Heat Release varied throughout the tests.
Samples B, C, and D had worse variability in their Heat Release than they did before the Air Distribution Tube.

## Slide 31: Questions or Comments?
Tina Emami
FAA William J. Hughes Technical Center
ANG-E212 Materials Fire Test
Building 275
Atlantic City Int’l Airport, NJ 08405
tina.emami@faa.gov
609-485-4277