## Slide 1: NexGen Burner for Insulation Burnthrough

## Slide 2: Introduction
Objective
To determine if the new ignitorless stator can be considered equivalent to the original stator baseline configuration for the NexGen burner when testing according to the insulation burnthrough rule 14CFR25.856b
Methodology
Both failure modes for the burnthrough rule will be evaluated for equivalency. These modes are:
  Burnthrough time: Polyacrylonitrile (PAN) material of two different areal weight will be tested in the picture frame sample holder, and the time to material burnthrough will be measured
    8579: 9 oz/yd2
    8611: 16 oz/yd2
  Backside Heat Flux: Nextel paper will be installed on the insulation burnthrough rig, and the backside heat flux will be measured with two Gardon heat flux gauges, one on each side of the test rig. The time to reach 2.0 BTU/ft2s will be used as the comparison point.
For both failure modes, the performance of the new stator at various inlet air pressures will be compared to the performance of the NexGen burner configured according to the parameters outlined in AC25.856-2A

## Slide 3
Configuration | Description
Baseline 60 psig | Original Stator with igniters
New Stator 60 psig | New Stator no igniters 60 psig air pressure
New Stator 65 psig | New Stator no igniters 65 psig air pressure
New Stator 70 psig | New Stator no igniters 70 psig air pressure
New Stator 75 psig | New Stator no igniters 75 psig air pressure
New Stator 80 psig | New Stator no igniters 80 psig air pressure
Description of various burner configurations for burnthrough testing
Baseline Stator Configuration (Left) and New Stator Configuration (Right)

## Slide 4: PAN Burnthrough Results
Chart: PAN Material Type
| 8579 | 8611
Baseline | 244.3 | 298.4
new stator, 60 psig | 249.5 | 318
new stator, 65 psig | 239.5 | 306
new stator, 70 psig | 224.5 | 294
new stator, 75 psig | 217 | 280.5
new stator, 80 psig | 230 | 296
Configuration | 8579 Burnthrough | 8611 Burnthrough
Baseline 60 psig | 244.3 | 298.4
New Stator 60 psig | 249.5 | 318.0
New Stator 65 psig | 239.5 | 306.0
New Stator 70 psig | 224.5 | 294.0
New Stator 75 psig | 217.0 | 280.5
New Stator 80 psig | 230.0 | 296.0
Table 2. PAN Material Burnthrough Results for Varying Burner Inlet Air Pressures
Figure 2. Comparison of PAN Burnthrough Time Against Burner Inlet Air Pressure
65 psig provides closest BT times to baseline original stator

## Slide 5: PAN Burnthrough Observations
A definite correlation is seen between increasing air pressure (and hence air velocity) and decreasing material burnthrough times for both PAN materials up to 75 psig.
Beyond 75 psig, the correlation reverses. The cause of this is unknown.
65 psig with the new stator gives the closest burnthrough times to the baseline for both PAN materials

## Slide 6: Backside Heat Flux Results
Time to Measured Heat Flux of 2.0 BTU/ft2s for Nextel Paper
Comparison of Measured Backside Heat Flux for Nextel Paper
Chart: Sonic Choke Inlet Pressure, psig
| 60 | 65 | 70 | 75 | 80
Baseline Left | 119 | 103 | 93 | 53 | 58
Baseline Right | 110 | 120 | 76 | 80 | 63
New Stator Left | 40 | 61 | 75 | 55 | 55
New Stator Right | 32 | 59 | 57 | 59 | 59
Two sheets of Nextel paper placed on burnthrough test rig
 |  | Time to 2.0 BTU/ft2s (seconds) |  |  |  | 
 | Pressure (psig) → | 60 | 65 | 70 | 75 | 80
Baseline | HFG Left | 119 | 103 | 93 | 53 | 58
 | HFG Right | 110 | 120 | 76 | 80 | 63
New Stator | HFG Left | 40 | 61 | 75 | 55 | 55
 | HFG Right | 32 | 59 | 57 | 59 | 59

## Slide 7: Backside Heat Flux Observations
A significant difference in backside heat flux measurement was found between the baseline configuration and new stator
  Although 65 psig provided the most similar PAN burnthrough times, it resulted in a heat flux failure 51 seconds earlier than the 60 psig baseline
An opposing correlation was found between heat flux failure time and sonic choke inlet pressure for the two configurations
  For the baseline configuration, as the sonic choke inlet pressure increases, the time to 2.0 decreases
  For the new stator configuration, as the sonic choke inlet pressure increases, the time to 2.0 initially increases, then levels off
Chart: Sonic Choke Inlet Pressure, psig
| 60 | 65 | 70 | 75 | 80
Baseline Left | 103 | 53 | 58
Baseline Right | 110 | 76
New Stator Left | 55 | 55
New Stator Right | 32 | 59 | 57
Δ=78s
Δ=44s
Δ=19s
Δ=2s
Δ=3s

## Slide 8: Heat Flux on Picture Frame
Single sheet of Nextel paper on picture frame holder with Gardon gauge centered on back and 12” away from back surface
Chart: Time, seconds
New stator pressure settings 65, 75, 80 psig are within 0.03 BTU/ft2s (1.5% F.S.)
New stator pressure setting 70 psig is within 0.06 BTU/ft2s (3% F.S.)
New stator pressure setting 60 psig is within 0.10 BTU/ft2s (5% F.S.)
Time (min) | 60 PSIG Baseline | 60 PSIG | Δ | 65 PSIG | Δ | 70 PSIG | Δ | 75 PSIG | Δ | 80 PSIG | Δ
1 | 2.74 | 2.65 | 0.09 | 2.75 | 0.01 | 2.81 | 0.07 | 2.80 | 0.06 | 2.73 | 0.01
2 | 2.84 | 2.76 | 0.08 | 2.76 | 0.08 | 2.90 | 0.06 | 2.82 | 0.02 | 2.80 | 0.04
3 | 2.90 | 2.81 | 0.09 | 2.93 | 0.03 | 2.94 | 0.04 | 2.90 | 0.00 | 2.89 | 0.01
4 | 2.93 | 2.78 | 0.15 | 2.94 | 0.02 | 2.99 | 0.07 | 2.90 | 0.02 | 2.99 | 0.07
 |  | Average Δ | 0.10 |  | 0.03 |  | 0.06 |  | 0.03 |  | 0.03

## Slide 9: Nextel on Picture Frame Observations
 | New Stator Pressure Setting |  |  |  | 
 | 60 | 65 | 70 | 75 | 80
Average Δ from Baseline | 0.10 | 0.03 | 0.06 | 0.03 | 0.03
% F.S. (2.0 BTU/ft2s) | 5.00 | 1.50 | 3.00 | 1.50 | 1.50
On the picture frame sample holder, all new stator pressure settings provide a measured heat flux similar to the baseline
  All within 5% full scale
The 65 psig setting, which was found to have the closest PAN burnthrough times to the baseline, provided a measured backside heat flux within 0.03 BTU/ft2s on average (1.5% of full scale)

## Slide 10: Insulation Blanket on Picture Frame
Time (min) | 60 PSIG Baseline | 60 PSIG | Δ | 65 PSIG | Δ | 70 PSIG | Δ | 75 PSIG | Δ | 80 PSIG | Δ
1 | 1.61 | 1.54 | 0.07 | 1.42 | 0.19 | 1.65 | 0.04 | 1.45 | 0.16 | 1.48 | 0.13
2 | 1.69 | 1.83 | 0.14 | 1.71 | 0.02 | 1.88 | 0.19 | 1.72 | 0.03 | 1.72 | 0.03
3 | 1.76 | 1.97 | 0.21 | 1.89 | 0.13 | 2.01 | 0.25 | 1.85 | 0.09 | 1.88 | 0.12
4 | 1.76 | 2.02 | 0.26 | 1.98 | 0.22 | 2.05 | 0.29 | 1.91 | 0.15 | 1.96 | 0.20
 |  | Average Δ | 0.17 |  | 0.14 |  | 0.19 |  | 0.11 |  | 0.12
Chart: Time, Minutes
.34 pcf FG
Metallized Tedlar
Metallized Tedlar
Nextel Paper
All new stator pressure settings are within 0.20 BTU/ft2s (10% F.S.)

## Slide 11: Insulation Blanket on Picture Frame Observations
On the picture frame sample holder, all new stator pressure settings provide a measured heat flux similar to the baseline
  All within 10% full scale
The 65 psig setting, which was found to have the closest PAN burnthrough times to the baseline, provided a measured backside heat flux within 0.14 BTU/ft2s on average (7% of full scale)
The addition of the film and fiberglass insulation have an influence on the measured heat flux
The flames created by the different configurations are not exactly the same, and have different effects on the insulation blanket, resulting in slightly differing backside measured heat flux
 | New Stator Pressure Setting |  |  |  | 
 | 60 | 65 | 70 | 75 | 80
Average Δ from Baseline | 0.17 | 0.14 | 0.19 | 0.11 | 0.12
% F.S. (2.0 BTU/ft2s) | 8.50 | 7.00 | 9.62 | 5.38 | 6.00

## Slide 12: Insulation Blankets on Test Rig
Chart: Time, minutes
 | 60 PSIG Baseline |  | 65 PSIG |  | Δ | 
Time (min) | Left | Right | Left | Right | ΔLeft | ΔRight
1 | 0.73 | 0.52 | 0.37 | 0.49 | 0.36 | 0.03
2 | 0.84 | 0.60 | 0.69 | 0.88 | 0.15 | 0.28
3 | 0.91 | 0.70 | 1.17 | 1.15 | 0.26 | 0.45
4 | 0.98 | 0.79 | 1.33 | 1.36 | 0.35 | 0.57
 |  |  |  | Average Δ | 0.28 | 0.33
.34 pcf FG
Metallized Tedlar
Metallized Tedlar
Nextel Paper
65 psig is within 0.30 BTU/ft2s (15% F.S.)

## Slide 13: Insulation Blanket on Test Rig Observations
On the test rig, the new stator configuration can be seen to provide a different heat flux profile during a test
The new stator configuration at 65 psig is within 0.3 BTU/ft2s of the baseline (15% of full scale)
The new stator configuration can be found to provide a more uniform flame (better left-to-right heat flux gauge agreement)
For this blanket construction, the 15% difference in heat flux measurements results in 4 minute heat flux of less than 2.0 BTU/ft2s for both burner configurations
 | 60 PSIG Baseline |  | 65 PSIG |  | Δ | 
Time (min) | Left | Right | Left | Right | ΔLeft | ΔRight
1 | 0.73 | 0.52 | 0.37 | 0.49 | 0.36 | 0.03
2 | 0.84 | 0.60 | 0.69 | 0.88 | 0.15 | 0.28
3 | 0.91 | 0.70 | 1.17 | 1.15 | 0.26 | 0.45
4 | 0.98 | 0.79 | 1.33 | 1.36 | 0.35 | 0.57
 |  |  |  | Average Δ | 0.28 | 0.33

## Slide 14: New Stator Testing Summary
The new stator was found to provide similar results to the baseline configuration when adjustments are made to the sonic choke inlet pressure
  PAN burnthrough times at 65 psig were within 2-3% of the baseline configuration
  At 65 psig, backside heat flux measurements were within 1.5% of full scale when Nextel paper was tested on the picture frame
  Testing an actual insulation blanket introduced variables that increased the variation in heat flux measurements between the baseline and new stator configurations
The new stator is favorable over the original stator for several reasons
  Ease of setup: No igniters, wires, etc. in the draft tube. Less dimensions to specify
  Better consistency: left-to-right heat flux measurements variations reduced with new setup

## Slide 15: Burnthrough Task Group
Meet to discuss new burner setup
Need input from task group
Need some “actual” insulation blankets to further investigate differences b/w baseline and new setup
Possible round robin with new stator setup

## Slide 16
Questions?
Robert I. Ochs
Fire Safety Branch
William J. Hughes Technical Center
ANG-E212; Bldg 287
Atlantic City, NJ 08405
T 609 485 4651
E robert.ochs@faa.gov