## Slide 1: A Statistical Approach to OSU Variability Reduction
Ben Grogan – BR&T Composites & Flammability
Shobbo Basu – Ph.D, ATF, BR&T Applied Math
Greg Hooker – BCA Process Engineering
Dan Slaton – TF, BCA Flammability

## Slide 2: Contents
Boeing OSU Variability Reduction Project
Statistical Evaluation of FAATC RR2012 Data
Recommendations

## Slide 3: Boeing OSU Variability Reduction Project
In April 2010, Boeing began a project to better understand OSU variability between three machines at two labs
The Boeing Flammability team identified 69 variables
Key variables included:
  Sample position w/in the burn chamber
  Machine insulation
  Calorimeter (surface damage, position w/in calibration fixture)
  Airflow through chamber
  For additional variables, click here
Initial approach focused on evaluating high impact, easy to change variables – still insufficient for scope of project
Each variable was then addressed through scientific Design of Experiments (DOE) with the help of Boeing Applied Math

## Slide 4: Boeing OSU Variability Reduction Project
The DOE Approach:
  What is a DOE
  Why do we do it?
  OSU Project Example

## Slide 5: Boeing OSU Variability Reduction Project - What is Design of Experiments (DOE) ?
Design of Experiments involves the active, systematic, and controlled change of process (or product) inputs to induce and observe their effects on process (or product) outputs.
A systematized collection of principles for valid experimentation (Randomization, Blocking, Replication and Balance)
A rigorous science for linking “how it was done” to “what you can say.”
A technology that includes techniques for structuring very powerful experiments with relatively few runs.

## Slide 6: Boeing OSU Variability Reduction Project - Experimentation Process
Prepare
Select
design
Conduct
experiment
Analyze
results
Verify
results
Document and
implement
Results
acceptable
?
More
experiments
?
Inadequate
preparation
Design
refinements
required
Errors in
execution
No
Yes
No
Optimize new
parameters
Add runs to
existing design
Yes

## Slide 7: Boeing OSU Variability Reduction Project - Example
Goal: Establish the thermopile baseline
Four critical variables.
  Lower Flame Length
    Nominal Length = Diameter (D) of the burner tube, 0.5D and 1.5D
    0.004 ft^3/min methane and an air supply adjusted to produce a flame such that the inner cone is approximately the same length as the diameter of the burner tube
  Upper Flame length
    Nominal 1”, also try 0.75” and 1.25”
    Controlled by Methane mixed with air in a ratio of approximately 50/50 by volume. Flamelets should be approximately one inch long (0.25" tip appears yellow)
  Heat flux (Center)
    Nominal (3.5 W/cm^2) and the upper/lower ranges ( + 0.05 W cm^2)
  Airflow
    Nominal (200mm Hg). Vary between 180mm Hg and 220mm Hg
Total combinations = 3 x 3 x 3 x 3 = 81
Based on DOE principles, adequate data can be collected through 48 combinations with multiple operators

## Slide 8: Boeing OSU Variability Reduction Project
Results, after many iterations of DOE’s:
  Understanding of impact of key variables, and how to control them
  Standard Work Instructions created to control ‘standardized’ parameters – Daily Startup Checklist, Maintenance Checklist, process instructions
  Checklists were developed based on weekly discussions on machine operation.
  Updated regularly to ensure consistent operation of machines with multiple users
  Standard Panel tested weekly to monitor and record progress on all machines – 5 panels per set
  Control charts maintained, regularly reviewed by non-advocate reviews
  Project continued for approximately one year
  Decreasing trend in standard deviations show reduced variability
  Ongoing testing to monitor OSU performance based on documented, standard procedures for setup, operation, and cleaning

## Slide 9: Boeing OSU Variability Reduction Project
Results:
  Daily startup checklist & maintenance checklist
  Have been provided to FAATC for inclusion in AC

## Slide 10: Boeing OSU Variability Reduction Project
Conclusion:
  Methodical approach successfully identified key parameters that contribute to variability
  DOE approach validated importance of standard work instructions, checklists, etc.
  Boeing’s OSU machines providing consistent, repeatable data
  Even with using established controls, there is “inherent” variability within the OSU
This method required up-front effort with data collection – many samples were tested over many months
BUT - the benefits from this approach were realized in a statistically significant body of data

## Slide 11: Applying Boeing’s Approach to the 2012 FAA Round Robin
Background:
  Test Approach
    36 OSU’s participated in testing
    Key equipment/setup information surveyed
    Testing conducted per individual lab standards
  Analysis Approach
    Results plotted to show side-by-side comparison (FAA)
    Operational variables statistically analyzed independently of test results (Boeing)
    Test results statistically analyzed to determine correlation between labs (Boeing)

## Slide 12: Applying Boeing’s Approach to the 2012 FAA Round Robin
Side-by-side plot shows a substantial body of data, with clearly variable data in both the test results and standard deviations
http://www.fire.tc.faa.gov/pdf/materials/June12Meeting/Burns-0612-OSU_HR2_Prototype_Data.pdf

## Slide 13: Applying Boeing’s Approach to the 2012 FAA Round Robin
Each lab operated based on its normal practices
Machine operating parameters were reported
Boeing analysis took operating parameters into account
Correlations drawn between each independent parameter
  Equipment manufacturer
  Inlet pressures
  Compressor vs. blower
  Etc.
Machine parameters then analyzed in reference to test data

## Slide 14: Applying Boeing’s Approach to the 2012 FAA Round Robin
Boeing equipment/setup analysis shows:
  Correlation between kW/mv and kW/mV*m^2 parameters (r=1, as expected)
  Strong positive correlation (r=0.8981) between flame off and flame on baselines (as expected)
  Negative correlation between calibration variables and baseline variables: a high Kh machine will have a lower baseline (r=-0.47/-0.58, not expected)
Notes: The closer the r value is to 1, the better the correlation (1/-1=perfect)

## Slide 15: Applying Boeing’s Approach to the 2012 FAA Round Robin
Conclusion:
  The statistical analysis identifies several equipment/setup parameters that significantly affect variability
    Machine Type
    Blower vs. Compressor
    Gas Pressure
  Not all variation is explained by observed parameters
    These factors explain less than 50% of the variability (R^2 ≤ 0.5)
  There are too many uncontrolled parameters and too little data to explain the variability
  Better control needs to be exercised over future Round Robin testing if variation is going to be explained

## Slide 16: Applying Boeing’s Approach to the 2012 FAA Round Robin
Recommendation:
  Conduct another focused Round Robin (5-10 machines)
    Much discussion needed in task group to define plan
    Cross section of machines – principle of balance
  Control parameters
    All labs set up to the same parameters: air/methane pressure, etc.
    Calibrate using the same process
    Calibrate using a single, shared calorimeter
    Collect a larger dataset – 10 to 50 panels tested at each lab
  Statistical analysis of collected data
  Use this standard approach in future development of the HR2

## Slide 17
Thank you for your time
Good luck OSU Task Group!